# M-Duino Trigger Timing Characterisation Guide

Detailed oscilloscope procedure with connection diagrams for delay, jitter, and synchronisation checks.

| Document date | 2026-05-02 |
| --- | --- |
| Application | M-Duino trigger coordination |
| Primary use | Deflagration trigger-path characterisation |
| Focus | Delay, jitter, relative skew, and connection setup |

# 1. Purpose

This guide explains how to measure the timing behaviour of the M-Duino when it is used as a trigger coordinator for a hydrogen deflagration setup. The objective is to quantify controller delay, shot-to-shot jitter, and the relative synchronisation offsets between the DAQ trigger branch and the camera trigger branch.

# 2. Connection Diagram

The figure below shows the recommended oscilloscope hookup logic for the three main measurement runs.

Use each panel as a separate measurement configuration. Do not try to measure everything at once if doing so makes the probing unclear or unreliable.

# 3. What This Characterisation Is Trying To Prove

The M-Duino is not the final scientific timing reference in this setup. The final scientific time base is provided by the high-speed datalogger. The main question is therefore whether the M-Duino distributes the trigger events with low enough uncertainty for the experiment.

Low and stable trigger delay

Low trigger jitter

Low relative skew between systems

# 4. Assumed Trigger Architecture

A clean input event is applied to the M-Duino.

The M-Duino produces a DAQ trigger output.

That DAQ trigger is split through a BNC splitter to DAQ 1 and DAQ 2.

The camera is triggered through its own separate trigger path.

Final scientific timestamping is performed by the high-speed datalogger.

# 5. Why Each Measurement Is Needed

Input to M-Duino DAQ trigger output: controller reaction delay and jitter.

Input to camera trigger path: controller-to-camera path timing and jitter.

M-Duino DAQ output to DAQ 1 and DAQ 2 inputs: splitter-path delay and branch-to-branch comparison.

Camera trigger path relative to the DAQ trigger path: inter-system synchronisation offset and stability.

# 6. Do You Need Special Benchmark Code?

Not necessarily for the first pass. A real-system timing study with the actual firmware and actual trigger paths is usually the most important measurement because it answers whether the experiment-ready system is good enough.

Special benchmark firmware of the type shown in the Industrial Shields article is optional and mainly useful if the real-system timing looks worse than expected or if the delay needs to be decomposed into hardware and firmware contributions.

# 7. Safety and Measurement Validity

Perform the test without ignition and without hazardous experiment execution.

Use a safe and repeatable pulse source.

Use the same firmware version intended for the experiments.

Use the same trigger edge that will be used experimentally.

Keep cable routing and terminations realistic.

Record all settings carefully.

Do not attach oscilloscope grounds carelessly. Only probe nodes that are safe for direct oscilloscope connection, and follow the lab grounding rules. If there is any doubt about common ground or floating circuits, use the correct differential or isolated measurement method.

# 8. Equipment Needed

M-Duino controller with the intended firmware

Oscilloscope, ideally 4 channels

Suitable probes

BNC cables

BNC splitter

Clean pulse source or function generator

Access to the chosen M-Duino input, the M-Duino DAQ trigger output, DAQ 1 input, DAQ 2 input, and the camera trigger line

# 9. How To Read The Three Diagram Runs

## 9.1 Run 1. Controller delay measurement

Use this setup to measure input to M-Duino DAQ trigger output and input to camera trigger path. It tells you the controller-side delay and jitter.

## 9.2 Run 2. DAQ splitter timing

Use this setup to measure M-Duino DAQ output to DAQ 1, M-Duino DAQ output to DAQ 2, and the relative skew between DAQ 1 and DAQ 2.

## 9.3 Run 3. Camera versus DAQ timing

Use this setup to measure the relative timing between the camera trigger branch and the DAQ trigger branch.

# 10. Exactly What To Measure On The Oscilloscope

Use the same edge definition everywhere. The best practice is to use the real experimental trigger edge and to measure timing at the 50 percent amplitude crossing.

t_input_to_daq = t(DAQ trigger output) - t(input event)

t_input_to_camera = t(camera trigger path) - t(input event)

t_daq1 = t(DAQ 1 trigger input) - t(M-Duino DAQ output)

t_daq2 = t(DAQ 2 trigger input) - t(M-Duino DAQ output)

skew_daq = t(DAQ 1 trigger input) - t(DAQ 2 trigger input)

skew_cam_daq1 = t(camera trigger path) - t(DAQ 1 trigger input)

skew_cam_daq2 = t(camera trigger path) - t(DAQ 2 trigger input)

# 11. Recommended Scope Configuration

Trigger on the input event for controller-delay measurements.

Trigger on the M-Duino DAQ output for branch comparison measurements.

Use appropriate vertical scale for the logic level.

Use enough time resolution to see the edges cleanly.

Use repeated acquisitions or statistics mode for jitter measurements.

Start with single acquisitions to confirm polarity and signal order.

# 12. Step-By-Step Measurement Procedure

Confirm that the pulse source produces a clean and repeatable edge.

Confirm that the pulse reaches the intended M-Duino input.

Confirm that the M-Duino produces the expected trigger outputs.

Confirm that the DAQ trigger line and camera trigger line are correctly identified.

Measure controller delay first, then splitter timing, then camera-versus-DAQ timing.

Repeat each measurement over many shots and record statistics.

# 13. How To Decide If The Measurement Is Valid

Signals are clean and repeatable.

The same edge definition is used everywhere.

The same reference channel is kept across repeated runs.

The same firmware and mode are used throughout.

The probed points are the real experiment nodes.

# 14. Common Mistakes To Avoid

Measuring the wrong edge.

Measuring before the real output stage instead of at the actual output node.

Measuring only one DAQ branch and assuming the second one is identical.

Forgetting that the camera is on a different trigger path.

Changing firmware during the test campaign.

Using arbitrary cursor points instead of a consistent crossing definition.

# 15. Results Table

| Path | Mean delay | Min | Max | Std dev | Pk-Pk jitter | Repetitions | Notes |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Input -> M-Duino DAQ output |  |  |  |  |  |  |  |
| Input -> M-Duino camera trigger path |  |  |  |  |  |  |  |
| M-Duino DAQ output -> DAQ 1 input |  |  |  |  |  |  |  |
| M-Duino DAQ output -> DAQ 2 input |  |  |  |  |  |  |  |
| DAQ 1 input -> DAQ 2 input |  |  |  |  |  |  |  |
| Camera trigger path -> DAQ 1 input |  |  |  |  |  |  |  |
| Camera trigger path -> DAQ 2 input |  |  |  |  |  |  |  |

# 16. Timing Budget Interpretation

For this application, the M-Duino is mainly a coordination layer, not the final scientific time base.

Total synchronisation uncertainty ~= controller jitter + branch skew + device trigger-response jitter

# 17. Test Record Template

| Date |  |
| --- | --- |
| Operator |  |
| Firmware version |  |
| M-Duino mode |  |
| Input channel used |  |
| DAQ trigger output used |  |
| Camera trigger path |  |
| Oscilloscope model |  |
| Probe type |  |
| Trigger edge |  |
| Sample rate |  |
| Repetitions |  |